- Equipment: PRC laser (Tunable CO2 Laser)

- Website for manual: http://www.prclaser.com/sts_series.shtml

- Operation procedure for beginner

Turn on procedure:

- Turn on the power switch on the wall

- Turn on the two tap water yellow switches on the wall

- Turn on the chiller (larger button: ON; smaller button: LOCAL)

- Turn on the red compressor (AUTO) until 80 psi

- Turn on the yellow cooler

- Open the four gas cylinders (limit: 20 psi)

- Turn on the key on the front panel of the laser

- Push the “laser on” button on the control panel

- When “H.V. Ready”, push the “H.V. On” button on the control panel

- Push “F5” on the front panel to turn off the “cold” status

- When “Laser ready”, push “Open Loop” button on the control panel

- Once the wavelength is fixed (grey metallic window), push the “Beam On” button on the control panel

- Once the power is fixed (with “commanded power percentage”), push the “Open Shutter” button on the control panel.

Turn off procedure:

- Turn off the shutter key control on the control panel

- Push all the buttons off on the control panel

- Turn off the chiller

- Switch off the tap water circulation

- Push the “Manual backfill” button on the front panel of the laser (until 760 psi)

- Turn off the red compressor (OFF)

- Turn off the yellow cooler

- Close all gas cylinders

- Switch off the power on the wall (alarm light outside off)

- Cover all the mirrors and lenses